Comparison of the Pullout Forces of Bioabsorbable Polylactide/Polyglycolide Screws (Biosorb and Lactosorb) and Tacks: A Study on the Stability of Fixation in Human Cadaver Parietal Bones

Johanna Tiainen1, Sanna-Mari Leinonen1, Jouko Ilomäki2, Esa Suokas2, Pertti Törmälä3, Timo Waris1, Nureddin Ashammakhi1,3

- Department of Surgery, Oulu University Hospital, Oulu, Finland.

- Linvatec Biomaterials Ltd., Tampere, Finland.

- Institute of Biomaterials, Tampere University of Technology, Tampere, Finland.

The aim of this study was to compare the pull-out forces of bioabsorbable polylactide/glycolyde (PLGA) tacks and screws in human cadaver parietal bones. Parietal bone pieces (c. 6 cm x 20 cm) were collected from five human male cadavers (age 47 to 75 years). Forty-nine BioSorbPDX® self-reinforced (SR-PLGA 80/20) tacks (diameter 1.5 mm, length 4.0 mm), 47 BioSorbPDX® (SR-PLGA 80/20) screws (diameter 1.5 mm, length 4.0 mm) and 46 LactoSorb® (PLGA 82/18) screws (diameter 1.5 mm, length 4.0 mm) were applied. The tacks were applied to drill holes using a special applicator gun (no tapping or tightening). The screws were applied to drill holes in the traditional way using tapping, and tightening with a screwdriver. A tensile testing machine was used. All the implants were tested thus: the head of the implant was held by an aluminum jig and the jig was pulled with wire until implant failure. The testing pull speed was 10 mm/min. Means and standard deviations (SDs) were calculated and the data analyzed using ANOVA. The pull-out force of the tacks was 115.9 ± 8.3 N, that of Lactosorb® screws was 112.9 ± 12.1 N and that of Biosorb® screws was 110.4 ± 8.9 N (statistically insignificant difference between the three groups). The most common reason for failure in the case of tacks was barb breakage (55.1%); it was thread breakage in the case of BioSorb® screws (66%) and stem split in the case of Lactosorb® screws (56%). Tacks seem to have a similar, perhaps even a little better holding power to cranial bone as screws and can hence be recommended for clinical application, as the procedure saves time and, consequently, costs. 

Keywords: Applicator gun, Bioabsorbable, Screw, SR-PLGA, Tack